## Slide 1: Clinical Principles of Intraosseous Vascular Access ARROW® EZ-IO® Intraosseous Vascular Access System
Notes: Thank you for your interest in the ARROW® EZ-IO® Intraosseous Vascular Access System. This program is designed to provide the practitioner with didactic knowledge and should be used in conjunction with hands-on simulation practice. If at any time you have a question that needs clarification, please contact us at 866-479-8500 or online.
Key concepts covered in this program include:
Indications and Contraindications
Basic bone anatomy
Site selection/identification
Needle set selection
Insertion technique
Infusing medications and fluids
Care of the patient
Assessment/documentation
Removing the catheter
Care of the driver
Clinical support resources

## Slide 2: Disclosure
Presenter Information & Disclosure as applicable
Notes: Paid Consultant of Vidacare LLC, a subsidiary of Teleflex Incorporated
or
Employee of Teleflex Incorporated

## Slide 3: Indications & Contraindications
The ARROW® EZ-IO® Intraosseous Vascular Access System is indicated for adult and pediatric patients any time vascular access is difficult to obtain in emergent, urgent or medically necessary situations for 24 hours
Contraindications
Fracture of target bone
Infection at area of insertion
Inability to identify landmarks
IO or attempted IO access in target bone within previous 48 hours
Prosthesis or orthopedic procedure near insertion site
Adults | Pediatrics
Proximal humerus Proximal tibia Distal tibia | Distal femur Proximal humerus Proximal tibia Distal tibia
Indications
Notes: The ARROW® EZ-IO® Intraosseous Vascular Access System is designed to be a safe, effective solution for fast vascular access when you’re facing difficult vascular access (DVA) challenges in emergent, urgent or medically necessary situations for adult and pediatric patients for 24 hours.
The system provides vascular access to the central circulation within seconds for the delivery of medications, intravenous (IV) fluids and blood products providing peripheral access with central performance.
There are few contraindications to IO access, and they are all relevant to site selection:
Trauma/Fracture (in the targeted bone) - If there is a fracture (or suspected fracture) in the bone in which the needle set is to be placed, an alternate site must be chosen. When IO access is placed in a fractured bone, fluid may extravasate into the surrounding tissue through the fracture. This can lead to complications including compartment syndrome. If a tibial fracture is suspected, both the proximal and distal sites are contraindicated as they share a common pathway inside the bone. However, if the patient has a fractured femur and an intact tibia, the tibial sites could be considered – circulation from the lower leg should be assessed to determine the appropriateness for IO vascular access.
IO or attempted IO access in the target bone within the past 48 hours: Healing from intraosseous insertion generally takes approximately 48 hours and is required before another IO catheter can be safely placed in the same bone. Within 48 hours fibrin formation and clotting are sufficient to prevent infiltration/extravasation through the previous IO access hole. If IO access was attempted in the bone, the same consideration applies. Complete healing, to the point where the hole can no longer be detected by Xray usually takes several days or weeks.1
Prosthesis or previous orthopedic procedures near insertion site: If the patient has a surgical scar over a joint; assume there is a titanium appliance within the joint – and do not attempt IO access at that site. Select a different insertion site and assess that site for use.
Infection at the insertion site; do not insert the needle set through infected or inflamed tissue.
Inability to locate landmarks (e.g. due to excessive tissue): Inability to locate insertion site landmarks for any reason (including excess adipose tissues) is a contraindication - select an alternate insertion site where landmarks are identifiable.
1 Miller LJ, Philbeck TE, Puga TA, Montez DF, Escobar GP. A pre-clinical study to determine the time to bone sealing and healing following intraosseous vascular access. Ann Emerg Med 2011;58(4S):S240. Teleflex Incorporated (Vidacare LLC) Sponsored Research

## Slide 4: Intraosseous Vascular Access History
1922
Drinker
1942
Papper
1945
WWII
1985 Orlowski
Global Leaders: Emergency & Critical Care
American Heart Association (AHA)
European Resuscitation Council (ERC)
International Liaison Committee on Resuscitation
(ILCOR)
1988 PALS
2005 ACLS
Notes: Intraosseous (IO) access has endured for more than 90 years as a safe and effective alternative for vascular access. The technique of intraosseous vascular access was first described in medical literature by Drinker et al in the 1920s.1 In 1942, Papper established that circulation times of intraosseous and intravenous fluids were essentially the same and that it was a viable option to other infusion techniques.2
During World War II, there are reports of IO vascular access in over 4,000 adults, primarily by medics to save victims of hemorrhagic trauma. The use of IO vascular access largely disappeared after that war, owing to the fact that there was not yet an organized pre-hospital and emergency care system, and the plastic intravenous catheter was increasing in popularity.
Re-discovered in 1985 - James Orlowski MD submitted an editorial titled: “My Kingdom for an Intravenous Line”3 which is seen as the genesis for renewed interest in use of IO access in pediatrics.
A resurgence of IO vascular access in pediatric patients was seen; the thinner, softer bone cortex made it easier to push a needle through. IO vascular access was established as a standard of care in Pediatric Advanced Life Support (PALS) in 1988.4 With the advent of IO access technology designed to access the harder, thicker bone cortex of the adult, intraosseous became an adult standard of care when the American Heart Association added IO vascular access to Advanced Cardiac Life Support (ACLS) in 2005.5
Global leaders in emergency care including the American Heart Association (AHA), the European Resuscitation Council (ERC)6 and the International Liaison Committee on Resuscitation (ILCOR)7 endorse the early use of intraosseous access. These renowned scientific groups recommend the IO route be considered if intravenous access is difficult or impossible.
1 Drinker CK, Drinker KR, Lund CC. The circulation in mammalian bone marrow. Am J Physiol 1922:62:1-92.
2 Papper EM. Bone marrow route for injection fluids and drugs into the general circulation. Anesthesiology. 1942;3:307-13.
3 Orlowski J. My kingdom for an intravenous line. Am J Dis Child. 1984;138:803.
4 Luck RP, Haines C, Mull C. Intraosseous access. J Emerg Med. 2010;39(4):468-475.
5 American Heart Association. American Heart Association guidelines for cardiopulmonary resuscitation and emergency cardiovascular care: part 7.2, management of cardiac arrest. Circulation. 2005;112(24 suppl):IV-58-IV-66.
6 European Resuscitation Council. European resuscitation guidelines. Resuscitation. 2005;6751:51-52. http://www.erc.edu/index.php. Accessed September 30, 2010.
7 International Liaison Committee on Resuscitation (ILCOR). International consensus of cardiopulmonary resuscitation and emergency cardiovascular care science with treatment recommendations: advanced life support. Resuscitation. 2005;67(pt 4):213-247.

## Slide 5: Organizational Support
ASTNA
AACN
ITLS
NAEMSP
SPN
ACEP
CoTCC
INS
ACS
ENA
Clinical Papers
Position Statements
Program Inclusion
Notes: Organizational support for intraosseous vascular access can be found in various forms including Clinical Papers, Position Statements, Course & Program Inclusions from organizations including the:
Air & Surface Transport Nurses Association (ASTNA)
American Association for Critical-Care Nurses (AACN)
American College of Emergency Physicians (ACEP)
American College of Surgeons (ACS)
American Heart Association (AHA)
Committee on Tactical Combat Casualty Care (CoTCC)
Consortium on Intraosseous Vascular Access in Healthcare Practice
Emergency Nurses Association (ENA)
European Resuscitation Council (ERC)
Infusion Nurses Society (INS)
International Liaison Committee On Resuscitation (ILCOR)
International Trauma Life Support (ITLS)
National Association of EMS Physicians (NAEMPS)
Society of Pediatric Nurses (SPN)
The National Association of EMS Physicians (NAEMSP) published a position statement in 2007 encouraging IO access availability in the pre-hospital setting.1 The Infusion Nurses Society (INS) issued a position paper in 2009 supporting IO access as an option for difficult vascular access in emergent and non-emergent situations. The paper formalizes the competent nurse’s role in insertion, maintenance and removal.2 The INS position paper is supported by the American Association of Critical Care Nurses (AACN)3 and the Emergency Nurses Association (ENA)4 . IO access techniques are taught in ENA’s ENPC and TNCC courses and included in their evidence based document titled “Emergency Nursing Resource for Difficult Intravenous Access”.
Physicians support IO access as well. A 2011 clinical policy statement released by the American College of Emergency Physicians (ACEP) states “IO access is a standard of emergency medicine practice and must be available in cases of difficult vascular access…non-physicians should be allowed to use IO expeditiously”5 The ARROW® EZ-IO® Vascular Access System’s instructional information can be found on the ACEP smartphone app.
Additionally, the military is an advocate for IO access. The Committee on Tactical Combat Casualty Care recommends IO access in their guidelines for military and tactical team use. The list goes on, including the Air & Surface Transport Nurses Association and the Society of Pediatric Nurses. Additionally, you will find IO vascular access in the American College of Surgeons Advanced Trauma Life Support (ATLS) trauma program, International Trauma Life Support (ITLS) and many other course curriculums designed to ensure optimal care is provided anytime life saving measures may be required. Many residency and nursing programs also include IO access in their curriculums.
1 Fowler R, Gallagher JD, Isaacs MS, et al. The role of intraosseous vascular access in the out-of-hospital environment [resource document to the NAEMSP position statement]. Prehosp Emerg Care. 2007;11:63-66.
2 Infusion Nurses Society. The role of the registered nurse in the insertion of intra-osseous access devices [position paper]. J Infus Nurs. 2009;32(4):187-188.
3 American Association of Critical-Care Nurses. American Association of Critical-Care Nurses endorses nurses’ expanded role in use of IO access devices. ww.aacn.org/wd/pressroom/docs/INSIOaccessrelease.pdf. June 21, 2010. Accessed September 30, 2010.
4 Emergency Nurses Association. ENA-supported statements. August 2009. http://www.ena.org. Accessed September 30, 2010
5 American College of Emergency Physicians. http://www.acep.org/Clinical---Practice-Management/Alternative-Methods-to-Vascular-Access-in-the-Emergency-Department/ Accessed January 22, 2014.

## Slide 6: Publications
Evidence Based Care
500+ Intraosseous Access Research Articles
190 Clinical Articles Specific to the ARROW® EZ-IO® System*
70+ Studies & Clinical Trials
4700+ Patients Studied
Widespread Utilization of ARROW® EZ-IO® Vascular Access
Estimated 2 million patients
50 Countries
*Effective January 2014, the EZ-IO® System was re-branded as the ARROW® EZ-IO® System
Notes: The body of research regarding intraosseous vascular access is extensive. The safety and efficacy of IO access has been studied in a variety of environments in several countries. Researchers have studied IO vascular access in academic settings, pre-hospital field settings, and in austere environments such as battlefield and disaster situations.
As of July 2014 there are:
More than 500 articles on intraosseous vascular access
More than 170 articles (in six languages) regarding use of the ARROW® EZ-IO® Vascular Access System* including approximately 60 articles describing case studies or clinical trials and over 4700 patients studied
The ARROW® EZ-IO® Vascular Access System has been the vascular access of choice in an estimated 2 million patients, and is currently used in over 50 countries.
*Effective January 2014, the EZ-IO® System was re-branded as the ARROW® EZ-IO® System.

## Slide 7: Who Can Utilize The ARROW® EZ-IO® Vascular Access System?
MEDICAL PROFESSIONALS
Physicians
Registered Nurses
Pre-Hospital Providers
Physicians Assistants
Advanced Practice Nurses
TYPICAL REQUIREMENTS
Policy/Protocol
Education
Competency
Practice
Notes: Patients with difficult vascular access (DVA) present a challenge throughout the pre-hospital and hospital settings including the ED, ICU, step-down, OR, L&D, Peds units, general medical floor, and other areas.
Any qualified healthcare professional can establish ARROW® EZ-IO® vascular access in seconds; Physicians, Physician Assistants; Advanced Practice Nurses and Registered Nurses as permitted by state regulations governing their practice. Placing an ARROW® EZ-IO® Needle Set to establish vascular access is a simple procedure with minimal training requirements. A qualified and trained registered nurse is permitted to place and manage IO devices, if it is determined by regulation, position statement or decision-making model to be within that professional’s scope of practice.
Based on a review performed in July 2014, there are no US states where registered nurses are prohibited expressly by statute from placing and managing intraosseous access and infusions. Additionally, each US state permits a licensed, trained and qualified EMS professional to place and use IO devices upon the order of a medical director.
In the US, each state has laws and regulations that govern the medical procedures licensed personnel may perform within their respective scope of licensure. These laws, regulations (for example State Practice Acts) and directives are occasionally modified as new medical technology becomes standard and accepted within the healthcare industry.
There is no official “certification” process unless mandated by an agency/organization, medical director or hospital. Establishing vascular access using this system is similar to an IV catheter in that specific training must occur in order to use the device safely and correctly. We offer comprehensive educational programs designed specifically with your facility and recommend completion prior to using the device. Contact your representative to schedule an inservice, and access online education information at www.teleflex.com/ezioeducation

## Slide 8: When Can The ARROW® EZ-IO® Vascular Access System Be Used?
Emergent, urgent or medically necessary situations
  Immediate need for medications or fluids
  Prevention of delays to vascular access during critical situations
  Difficult Vascular Access (DVA)
  Vein preservation
Safe: <1% serious complication rate
Fast: Vascular access in seconds
Efficient: 97% first-attempt success rate
Versatile: Can be placed by any qualified healthcare provider
Convenient: Requires no additional equipment or resources
Notes: The ARROW® EZ-IO® Vascular Access System is safe, fast and easy, and can be used in any situation where vascular access is challenging. Research shows a serious complication rate of < 1%1 and clinicians have a 97%2 insertion success rate.
Placement of peripheral IV catheters for vascular access is common; however, in certain circumstances successful cannulation can be difficult and may require multiple attempts.3-5 Vascular access fails during emergency situations up to 40% of the time.6
Consider emergencies like hemodynamic instability and respiratory compromise. Outcomes in cardiac or respiratory arrest depend on immediate access. The patient that needs rapid sequence induction/intubation (RSI) may not have time to wait for other forms of vascular access. Patients with DVA may be in need of immediate medications or fluids: shock/trauma, sepsis, stroke, dysrhythmia/chest pain/myocardial infarction (MI), burns, status, epilepticus/seizures, diabetic ketoacidosis (DKA), end-stage renal disease (ESRD), head injury, anaphylaxis, congestive heart failure (CHF), sickle cell crisis/anemia and many other conditions are time sensitive requiring vascular access.
Any patient may present with difficult vascular access (DVA). With the ARROW® EZ-IO® Vascular Access System, vascular access is no longer a problem.
1 Rogers JJ, Fox M. The safety of intraosseous vascular access. Emergency Medicine Patient Safety Foundation Forum. Fall 2012:18-21 Teleflex Incorporated (Vidacare LLC) Sponsored Research
2 Cooper BR, Mahoney PF, Hodgetts TJ, Mellor A. Intra-osseous access (EZIO®) for resuscitation: UK military combat experience. JR Army Med Corps 2008;153(4):314-6.
3 Aponte H, Acosta S, Rigamonti D, Sylvia B, Austin P, Samolitis T. The use of ultrasound for placement of intravenous catheters. AANA J. 2007;75(3):212-216.
4 Bauman M, Braude D, Crandall C. Ultrasound-guidance vs. standard technique in difficult vascular access patients by ED technicians. Am J Emerg Med. 2009;27(2):135-140. 3
5 Costantino TG, Parikh AK, Satz WA, Fojtik JP. Ultrasonography-guided peripheral intravenous access versus traditional approaches in patients with difficult intravenous access. Ann Emerg Med. 2005;46(5):456-461.
6 Leidel BA, Kirchhoff C, Bogner V, et al. Is the intraosseous access route fast and efficacious compared to conventional central venous catheterization in adult patients under resuscitation in the emergency department? A prospective observational pilot study. Patient Saf Surg. 2009;3(1):1-8.

## Slide 9: Difficult Vascular Access Options
Right Line
Right Patient
Right Time
PIV
USGPIV
EJ
PICC
CVC
IO
Time
Risk
Cost
Skill
Notes: Intraosseous (IO) access technology is used as a low risk option for patients with DVA in emergent, urgent and medically necessary cases, and can be used in lieu of other options that are time consuming, skill intensive, higher risk and costly.1-2 The average time needed for peripheral IV catheterization is reportedly as high as 16 minutes in patients with DVA.3 Placing vascular access using the ARROW® EZ-IO® System can be achieved in seconds4 and eliminates the potential for multiple IV sticks, facilitating vein preservation.
Choosing the ARROW® EZ-IO® System for vascular access can expedite treatment and prevent departmental throughput delays, replacing time-consuming and skill intensive ultrasound guided peripheral IVs and eliminating external jugular lines that are risky, positional and not well tolerated by patients.1
Use of IO access technology reduces risk associated with treatment delays and vascular access therapies with higher complication rates.1,3 IO vascular access can be used as a bridge to the proper placement of a central venous catheter (CVC) and replace CVCs placed solely for vascular access.
ARROW® EZ-IO® Vascular Access is easily placed to treat patients (e.g. hypo- or hyperglycemia, hypovolemia, kidney stones, to provide thrombolytics, moderate sedation, analgesia, antibiotic therapy, for elective surgeries or as a rescue line in the unstable patient) in less time with less staff, less equipment and at a lower cost than many other vascular access options.5-8
Rapid access can halt the downward spiral that can occur during time consuming attempts to obtain vascular access, potentially reducing patient deterioration which can lead to rapid response or code situations.1
1 Phillips L, Brown L, Campbell T, Miller J, Proehl J, Youngberg B. Recommendations for the Use of Intraosseous Vascular Access for Emergent and Nonemergent Situations in Various Health Care Settings: A Consensus Paper. Journal of Infusion Nursing, November/December 2010; 33(6):346-351.
2 Emergency Nursing Resource: Difficult Intravenous Access. ENA; December 2012
3 Leidel BA, Kirchhoff C, Bogner V, Braunstein V, Biberthaler P, Kanz K-G. Comparison of intraosseous versus central venous vascular access in adults under resuscitation in the emergency department with inaccessible peripheral veins. Resuscitation. 2012;83(1):40-45.
4 Paxton JH, Knuth TE, Klausner HA. Proximal humerus intraosseous infusion: a preferred emergency venous access. J Trauma 2009;67:606–11.*
5 Dolister M, Miller S, Borron S, et al. Intraosseous vascular access is safe, effective and costs less than central venous catheters for patients in the hospital setting. J Vasc Access 2012; doi: 10.5301/jva.5000130.*
6 Fowler RL , Pierce A, Nazeer S et al. 1,199 case series: Powered intraosseous insertion provides safe and effective vascular access for emergency patients. Ann Emerg Med 2008;52:S152.
7 American Heart Association 2010 American Heart Association guidelines for CPR and ECC . Part 14-Pediatric Advanced Life Support. Circulation 2010;122(3 suppl.):S881-5.
8 Horton MA, Beamer C. Powered intraosseous insertion provides safe and effective vascular access for pediatric emergency patients. Pediatr Emerg Care 2008;24(6):347-50.*
*Teleflex Incorporated (Vidacare LLC) Sponsored Research

## Slide 10: Anatomy and Physiology
Highly vascular, non-collapsible access
Anatomy and Physiology
Rapid flush to displace marrow
Notes: IO catheters are usually placed in the proximal and distal ends (epiphyses) of long bones due to the thinner cortex and abundance of cancellous (spongy) bone at these sites. Within the medullary space lies a vast system of blood vessels through which blood and fluid pass into the central circulation. The bone acts as a large, non-collapsible vein. Fluid or medication instilled into the intraosseous space gains access to the central circulation within just a few seconds.1
The needle set is powered by the driver through bone that is usually 3mm thick or less; accurate identification of the insertion site is important to ensure you are not attempting access through thicker bone. The specially designed cutting IO needle set and powered driver enables clinician control of the insertion with proven tactile feedback.2 This is integral to success with newborns, infants and young children who may have thinner, softer bones. The ARROW® EZ-IO® Vascular Access System provides rapid, smooth entry into the bone’s medullary cavity, creating a conduit to the central circulation.
Within the micro-vasculature inside the medullary space lies a hyper-coagulable fibrin mesh as well as red and yellow marrow. The IO space is like a honeycomb of bone. The fibrin mesh and bone marrow is a thick substance and poses more resistance to infusion than blood. To achieve optimal fluid volume and drug delivery times, as well as less pressure for infusions, a rapid flush is necessary to displace the thick bone marrow.
1 Von Hoff DD , Kuhn JG, Burris HA, Miller LJ. Does intraosseous equal intravenous? A pharmacokinetic study. Amer J Emerg Med 2008;26:31-8.
2 Miller L, Philbeck T, Bolleter S, Garcia G. Tactile feedback comparison of three types of intraosseous access devices for needle insertion accuracy. Ann Emerg Med 2010;56(3):S133. Teleflex Incorporated (Vidacare LLC) Sponsored Research

## Slide 11: Real-time Fluoroscopy – Human Model
Notes: A common misperception is that it takes a long time for fluids, medications, or blood infused into the intraosseous space to reach the central circulation, when in fact medications and fluids instilled into the intraosseous space gain access to the central circulation within seconds. A preclinical study measuring peak serum concentrations found that epinephrine infused via the IO humeral site has the identical peak serum concentration as central line administration.1
This video of real-time fluoroscopy in a human subject demonstrates how quickly IV contrast infused into the medullary space is disseminated into the central circulation. Watch how quickly the dye reaches the right atrium when instilled via the proximal humerus, again demonstrating direct correlation between the subclavian and proximal humerus for infusion of fluids. The distance to the right atrium from the humerus is short; medication and fluid will reach the heart in less than 3 seconds.2 Typical fluid volume to expect from the proximal humerus site is 5L/hr under pressure.1 From the tibia to the heart there is more distance to travel, and typical fluid volumes of 1L/hr are achieved under pressure.3
1 Hoskins SL, Kramer GC, Stephens CT, Zachariah BS. Efficacy of epinephrine delivery via the intraosseous humeral head route during CPR. Circulation 2006;114:II_1204.*
2 Internal study report. Protocol 2013-06: Clinical Studies to Determine the Optimal Technique to Identify the Proximal Humerus Intraosseous Vascular Access Insertion Site. Vidacare Corporation, May 2013.
3 Philbeck TE, Miller LJ, Montez D, Puga T. Hurt so good; easing IO pain and pressure. JEMS 2010;35(9):58-69.*
*Teleflex Incorporated (Vidacare LLC) Sponsored Research

## Slide 12: Site Selection
Flow rates average 5L/hr
3 seconds to heart with medication/fluids
Lower insertion & infusion pain
Less medication required for pain management
No reported compartment syndrome due to IO access placement
High success rates in pediatric patients
  Do NOT use the powered ARROW® EZ-IO® Vascular Access System in the sternum!
Notes:   The ARROW® EZ-IO® Vascular Access System is indicated for adult insertion in 6 sites: the right and left proximal humerus, the right and left proximal tibia, and the right and left distal tibia. Indications for pediatric insertion include 8 sites: the right and left distal femur, the right and left proximal humerus, the right and left proximal tibia and the right and left distal tibia. The powered ARROW® EZ-IO® Vascular Access System includes the driver and needle sets and is not intended for placement in the sternum.
Site selection is based on several factors including patient size, anatomy, presenting condition, ability to locate anatomical landmarks, clinical judgment and experience. Many studies support the proximal humerus as a successful or even superior IO route for flow rates, drug delivery, and management of infusion pain.1-8 The proximity of the humerus to the heart ensures rapid infusion of medications into the central circulation. 5L/hr average flow rates can be achieved using the humeral site.3-5 Medications and fluids reach the right atrium in just 3 seconds. Research reports insertion into the proximal humerus results in lower insertion and infusion pain than the tibia and requires less medication for pain management.5 As of July 2014, there has been no reported compartment syndrome due to IO access placement at the proximal humerus.
The distal femur insertion site in pediatrics is supported in the medical literature as a viable site with success rates of 98-100%.9, 10
1 Kovar J, Gillum L. Alternate route: the humerus bone—a viable option for IO access. JEMS 2010;35(8):52-9.*
2 Wampler D, Schwartz D, Shumaker J, et al. Paramedics successfully perform humeral EZ-IO® intraosseous access in adult out-of-hospital cardiac arrest patients. Am J Emerg Med 2010;56(3):S88.VS doi.org/10.1016/j.ajem.2011.07.010.*
3 Miller L, Philbeck T, Montez D, Puga T. A two-phase study of fluid administration measurement during intraosseous infusion. Ann Emerg Med 2010;56(3):S151.*
4 Ong MEH, Chan YH, Oh JJ, Ngo AS-Y. An observational, prospective study comparing tibial and humeral intraosseous access using the EZ-IO®. Amer J Emerg Med 2009;27:8-15.*
5 Philbeck TE, Miller LJ, Montez D, Puga T. Hurts so good; easing IO pain and pressure. JEMS 2010;35(9):58-69*
6 Paxton JH, Knuth TE, Klausner HA. Proximal humerus intraosseous infusion: a preferred emergency venous access. The J Trauma 2009;67(3):606-11.*
7 Hoskins SL, Kramer GC , Stephens CT, Zachariah BS. Efficacy of epinephrine delivery via the intraosseous humeral head route during CPR. Circulation 2006;114:II_1204.*
8 Hoskins SL, Zachariah BS, Copper N, Kramer GC . Comparison of intraosseous proximal humerus and sternal routes for drug delivery during CPR. Circulation 2007;116:II_933.*
9 Horton MA, Beamer C. Powered intraosseous insertion provides safe and effective vascular access for pediatric emergency patients. Pediatric Emergency Care 2008;24(6):347-50.*
10 Dolister M, Miller ST, Borron S, Truemper E, Sah MR. Intraosseous vascular access can be used safely and effectively, and at a lower cost than central venous catheters, for pediatric and adult patients in the hospital setting. Ann Emerg Med 2011;58(4S):S311.*
* Teleflex Incorporated (Vidacare LLC) Sponsored Research

## Slide 13: Proximal Humerus Site Identification
Notes: Let’s watch this demonstration of correct identification of the proximal humerus site.
Proximal Humerus Insertion Site Identification
Place the patient’s hand over the abdomen (elbow adducted and humerus internally rotated)
Place your palm on the patient’s shoulder anteriorly
The area that feels like a “ball” under your palm is the general target area
You should be able to feel this ball, even on obese patients, by pushing deeply
Place the ulnar aspect of your hand vertically over the axilla and the ulnar aspect of your other hand along the midline of the upper arm laterally
Place your thumbs together over the arm
This identifies the vertical line of insertion on the proximal humerus
Palpate deeply up the humerus to the surgical neck
This may feel like a golf ball on a tee – the spot where the “ball” meets the “tee” is the surgical neck
The insertion site is 1 to 2 cm above the surgical neck, on the most prominent aspect of the greater tubercle

## Slide 14: Insertion Angle
Tibia and Femur: Insert needle set at a 90o angle to the bone
Humerus: Insert needle set at a 45o angle to the anterior plane and posteromedial
Notes: Insertion angle is important to ensure placement that will provide optimal vascular access.
For the proximal humerus insertion, aim the needle set tip at a 45-degree angle to the anterior plane and posteromedial.
For tibia and femur access, aim the needle set tip at a 90-degree angle to the bone.

## Slide 15: Proximal Tibia Site Identification
Adult
15
Infant/Child
Growth
Plate
Notes: Proximal Tibia Insertion Site Identification – Adult
Extend the leg. Insertion site is approximately 2 cm medial to the tibial tuberosity, or approximately 3 cm below the patella and approximately 2 cm medial, along the flat aspect of the tibia.
Proximal Tibia Insertion Site Identification – Infant/Child
Extend the leg. Pinch the tibia between your fingers to identify the medial and lateral borders of the tibia.
Insertion site is approximately 1 cm medial to the tibial tuberosity, or just below the patella (approximately 1 cm or one finger width) and slightly medial (approximately 1 cm), along the flat aspect of the tibia.

## Slide 16: Distal Tibia Site Identification
16
16
Midline
on the bone
Notes: Distal Tibia Insertion Site Identification - Adult
Insertion site is located approximately 3 cm proximal to the most prominent aspect of the medial malleolus. Palpate the anterior and posterior borders of the tibia to assure that your insertion site is on the flat center aspect of the bone.
Distal Tibia Insertion Site Identification – Infant/Child
Insertion site is located approximately 1-2 cm proximal to the most prominent aspect of the medial malleolus. Palpate the anterior and posterior borders of the tibia to assure that your insertion site is on the flat center aspect of the bone.

## Slide 17: Distal Femur Site Identification - Pediatrics
Secure the leg out-stretched
Ensure the knee does not bend
Notes: Distal Femur Insertion Site Identification - Infant/Child
Secure the leg out-stretched to ensure the knee does not bend. Identify the patella by palpation.
The insertion site is just proximal to the patella (maximum 1 cm) and approximately 1-2 cm medial to midline.

## Slide 18: ARROW® EZ-IO® Needle Set Selection
18
15 mm
15 gauge
Indicated for patients weighing 3-39 kg
25 mm
15 gauge
Indicated for patients weighing 3 kg or over
45 mm
15 gauge
Indicated for patients weighing 40 kg or over, excessive tissue depth
Clinical judgment should be used to determine appropriate needle set selection based on patient weight, anatomy and tissue depth overlying the insertion site
Notes: The single use sterile needle sets are 15 gauge, 304 stainless steel available in 3 lengths:
15 mm (pink hub) indicated for patients weighing 3-39 kg
25 mm (blue hub) indicated for patients weighing 3 kg or over
45 mm (yellow hub) indicated for patients weighing 40 kg or over and/or excessive tissue depth
The needle sets do not have “adult” or “pediatric” sizes. Each needle set is US FDA-cleared with weight range guidelines. Clinical judgment should be used to determine appropriate needle set selection based on patient weight, anatomy and tissue depth (skin, adipose and muscle thickness) overlying the insertion site.
For example, a small geriatric female may require a shorter length needle set; an obese child may require a longer needle set. In adults, the 45 mm needle set should be considered for the proximal humerus site in most patients weighing >40 kgs, to accommodate for any inadvertent movement of the extremity after insertion.

## Slide 19: Needle Set Selection
Estimate tissue depth
Confirm with 5 mm mark visible above the skin
Notes: To facilitate selection of the correct length needle set, tissue depth over the insertion site should be assessed prior to selection of the needle set. Press your thumb or finger over the insertion site to estimate the depth of soft tissue.
The catheter is marked with a black line 5mm from the hub. If the needle set is inserted through the soft tissue and does not reach the bone or the 5mm mark is not visible above the skin with the tip of the needle set touching the bone, the needle set is too short. A longer needle set or alternate site should be chosen prior to penetration of the bone cortex. Using a needle set that is too short will increase the risk of catheter dislodgement leading to infiltration/extravasation, creating a hole in the bone unnecessarily and rendering the site unusable for future IO access for a minimum of 48 hours.
Clinical experience with the device will ultimately present a more rapid approach to needle set selection, but the 5 mm mark assists the clinician with establishing which needle set is appropriate for the patient.
There is often concern expressed by clinicians when selecting a needle set that is long enough to enter the medullary space in infants and small children. With any insertion site or needle set length, clinicians learn to stop by releasing the trigger when they feel the loss of resistance as the needle set enters the medullary space. A bench study, with models designed to mimic bone density, has demonstrated the ease of identifying this “stopping point” for IO access insertion through tactile feedback.1
1 Miller L, Philbeck T, Bolleter S, Garcia G. Tactile feedback comparison of three types of intraosseous access devices for needle insertion accuracy. Ann Emerg Med 2010;56(3):S133. Teleflex Incorporated (Vidacare LLC) Sponsored Research

## Slide 20: Technique
Precision
Control
Gentleness
Speed
Cross-section of bone
Notes: This video portrays the precision, control and gentleness of this system utilizing a raw egg for demonstration. Allow the driver and the design of the needle set tip, patented to drill through bone, to do all the work. Insertion technique is gentle, reducing the risk for over-penetration, and produces a hole in the bone that is exactly the size of the catheter, reducing the risk of infiltration/extravasation.
With softer, smaller and thinner bones, like those in pediatric patients, special care must be taken during insertion to avoid excessive pressure and recoil. Recoil can occur when the clinician feels the lack of resistance upon entry into the medullary space, releases the trigger and inadvertently pulls back on the driver. This recoil may displace the needle set from the medullary space. We encourage you to practice insertion on a raw egg as part of the hands on component of training. A raw egg is a great simulation tool to help you learn to avoid recoil when entering the medullary space and releasing the trigger. Do NOT pull back on the driver when releasing the trigger.

## Slide 21: Prepare Supplies
Prep site
Prepare supplies
  Open EZ-Stabilizer® Dressing
  Prime the EZ-Connect® Extension Set
  Attach Needle Set to Driver
Notes: Use a clean, no-touch technique for this procedure: sterile equipment, decontaminated hands, non-sterile gloves, and avoidance of gloved hands or non-sterile surfaces contacting exposed sterile equipment or site that must remain uncontaminated to maintain asepsis. If using the distal femur site in pediatrics, secure the leg outstretched. Prep the site by using antiseptic solution of your choice.
Open the appropriate needle pack and remove the EZ-Connect® Extension Set.* Always prime the extension set with fluid before attaching to the catheter hub. The EZ-Connect® Extension Set primes with approximately 1.0 mL. The extension set has a luer lock that will twist onto the catheter hub.
If the patient is unresponsive to pain, attach a saline flush to the extension set and prime the tubing with saline, leaving the syringe and remaining saline attached to the extension set.
If the patient is responsive to pain, consider priming the extension set with 2% epinephrine and preservative-free intravenous lidocaine without for anesthesia prior to initial saline flush.
Open the EZ-Stabilizer® Dressing – the dressing is made with a strong medical adhesive tape that is simple to apply to secure the catheter. It is designed for the 15 mm, 25 mm and 45 mm needle sets.
Hold the driver in your dominant hand and attach the needle set by aligning the magnetized needle set hub to the driver.
Remove the safety cap from the needle set using standard sharps precautions. Stabilize the extremity.
*The EZ-Connect® Extension Set uses the Robertsite® needleless connector; the connector valve satisfies all CDC requirements for needleless intravascular systems.
Note: The EZ-Connect® Extension Set should NOT be used for high pressure infusion/power injection.

## Slide 22: Insertion Technique
Notes: With all patients, minimal pressure should be placed on the driver during insertion. Allow needle set rotation and gentle downward pressure to penetrate the bone.
Insertion Technique
Push the needle set tip through the skin until the tip rests against the bone
  Proximal Humerus - Aim the needle set tip at a 45-degree angle to the anterior plane and posteromedial
  Tibial and Femur Insertion - Insert needle set at a 90-degree angle to the bone
Check for the 5 mm mark – it must be visible above the skin for confirmation of adequate needle set length prior to drilling
Gently drill approximately 1-2 cm or until the hub is close to the skin in an adult; stop when you feel the loss of resistance in infants and small children.
  Apply the minimal amount of pressure required to keep the driver advancing into the bone. The needle set is advanced into the bone by the driver and patented needle tip, little pressure is required. Maintain light steady pressure for an adult insertion to maximize the speed of insertion. With pediatric patients, minimal pressure is required for successful placement. Insertion time varies as bone density varies from patient to patient.
Hold the hub in place and pull the driver straight off
NOTE: If the driver stalls and will not penetrate the bone, you may be applying too much downward pressure. In the unlikely event that the battery on the driver fails, clinicians may manually insert the needle set by removing the driver and grasping the needle set hub by hand – advance the needle set into the medullary space by penetrating the bone cortex with twisting, steady, firm pressure. Do NOT use excessive force, do NOT rock or bend needle set during insertion.

## Slide 23: Insertion Completion
23
Stabilize hub and remove driver
Remove stylet
Apply EZ-Stabilizer® Dressing
Firmly attach primed EZ-Connect® Extension Set
Place wrist band
23
Notes: To complete the insertion:
  Continue to hold the hub while twisting the stylet off the hub with counter clockwise rotations
  The catheter should feel firmly seated in the bone (1st confirmation of placement)
Place the stylet in a sharps container; always dispose of all sharps and biohazard materials using standard biohazard practices and disposal containers. If using the NeedleVISE® 1 port sharps block, place on stable surface and use a one-handed technique.
Place the EZ-Stabilizer® Dressing over the hub
Attach a primed EZ-Connect® Extension Set to the hub, firmly secure by twisting clockwise
Pull the tabs off the dressing to expose the adhesive, apply to the skin
Aspirate for blood/bone marrow (2nd confirmation of placement)
  Ensure clamp on extension set is open. Inability to withdraw/aspirate blood from the catheter hub does not mean the insertion was unsuccessful. Consider attempting to aspirate after the flush. Site placement can also be confirmed by ability to administer pressurized fluids, and noting the pharmacologic effects of medication administration after flow is established.
For IO access in the proximal humerus, secure the arm in place across the abdomen
  Due to the mobile nature and design of the humeral joint, arm movement and contraction of the deltoid muscle (e.g. simple abduction, rotation, or extension of the arm) can dislodge the IO catheter. Raising the arm horizontally above the plane of the torso will often impinge the IO catheter against the acromion and dislodge or bend the catheter. The arm should also be stabilized to limit movement.
For IO access in the distal femur (pediatrics), the leg should remain immobilized until the IO catheter is removed.
Be sure to apply the wrist band, which includes a place to indicate the date and time of placement as well as the 1-800 number to reach a clinician experienced with the ARROW® EZ-IO® Vascular Access System.

## Slide 24: Rapid Normal Saline Flush
24
Consider for blood typing and other commonly ordered labs
Infants & Children: 2-5 mL
Adults: 5-10 mL
Notes: Flush
It is essential to perform a rapid normal saline (NS) syringe flush into the IO space before attempting to infuse fluids through the catheter. A rapid syringe flush of 5-10 mL normal saline in adults and 2-5 mL normal saline in infants and children helps clear the marrow and fibrin from the medullary space, allowing for effective infusion rates. IO access may be compromised if the line is not used for prolonged periods. IO lines can often be opened successfully by an additional syringe flush. As with an IV infusion, an IO line should be assessed for patency and flushed before and after infusion. It is important to avoid extreme pressure for the flush, which may increase the risk of infiltration/extravasation.
Labs
Based on preclinical and clinical evidence comparing IO and venous or arterial sources a number of common laboratory values correlate well; other values show clinical similarity without statistically significant correlation, therefore caution should be exercised with their interpretation. Certain point of care analyzers have been studied with acceptable results. Check with your laboratory for IO specimen processing capabilities. For more information regarding IO specimen lab analysis, refer to the publication Science and Fundamentals of Intraosseous Vascular Access, available at: www.teleflex.com/ezioeducation
The following recommendations have been developed based on research done by Teleflex Incorporated (Vidacare LLC); study data was based on IO blood specimens obtained prior to any infusions or flush:
Connect a syringe directly to the catheter hub.
The first 2 mL of IO blood aspirate may be discarded or considered for point of care testing.
Samples must be identified as IO blood.

## Slide 25: Infusion & Medications
For optimal flow infuse with pressure
Administer medications in same dose, rate and concentration as given via peripheral IV
Notes: Infusions
As with other vascular access lines, IO access flow rates will vary among patients and anatomical sites. Generally, the proximal humerus allows higher flow rates than tibial sites. Adequate flow rates are dependent on performing a rapid flush (syringe bolus) prior to IO infusion and infusing fluids and medications under pressure (e.g. infusion pressure pump or pressure bag). Gravity alone will rarely generate adequate flow rates.
Pressure is the rate-limiting factor in achieving adequate flow rates in IO infusions; the higher the pressure, the greater the flow rate. Many infusion pumps (including rapid infusers such as the Level 1® ) are designed to administer large amounts of fluid very rapidly, but often work by measuring volume rather than infusion pressure. If adequate IO flow rates cannot be achieved with an infusion pump, a pressure bag should be considered.
Medications
Virtually any medication that can be given via a peripheral IV can be given via the IO route using the same dose, rate and concentration.1-4 Verify placement/patency prior to all infusions. Use caution when infusing hypertonic solutions, chemotherapeutic agents, or vesicant drugs. Follow medication administration with a flush.
1 Von Hoff DD , Kuhn JG, Burris HA, Miller LJ. Does intraosseous equal intravenous? A pharmacokinetic study. Amer J Emerg Med 2008;26:31-8.
2 Hoskins SL, Kramer GC , Stephens CT, Zachariah BS. Efficacy of epinephrine delivery via the intraosseous humeral head route during CPR. Circulation 2006;114:II_1204.*
3 Hoskins S, Nascimento P, Espana J, Kramer G. Pharmacokinetics of intraosseous drug delivery during CPR. Shock 2005;23:35.
4 Hoskins SL, Nascimento P Jr., Lima RM, Espana-Tenorio, JM, Kramer GC . Pharmacokinetics of intraosseous and central venous drug delivery during cardiopulmonary resuscitation. Resuscitation 2012;83(1):107-12.*
*Teleflex Incorporated (Vidacare LLC) Sponsored Research

## Slide 26: Pain Management
Slowly infuse initial dose over 120 seconds
Prime extension set with 2% lidocaine (IV lidocaine)
Allow lidocaine to dwell in IO space 60 seconds
Flush with normal saline
Slowly infuse 1/2 of initial dose over 60 seconds
Observe cautions/contraindications for lidocaine, confirm dose per institution
Notes: For patients able to perceive pain, there can be significant discomfort associated with the initial flush as well as infusion. Pressure changes occur within the bone as a result of infusion of medications and fluids, and if not properly anesthetized, pain associated with IO infusion under pressure is often severe. Anesthetizing the IO space prior to the initial flush should be considered for any patient who is conscious/alert to pain. If time constraints do not permit anesthetization prior to the initial flush, pain management should be considered as soon as time permits, and as needed throughout the course of IO infusion.
Recommended anesthetic for adult patients responsive to pain:
Observe recommended cautions/contraindications to using 2% preservative and epinephrine free lidocaine (intravenous lidocaine)
Confirm lidocaine dose per institutional protocol
Prime extension set with lidocaine Note that the priming volume of the EZ-Connect® Extension Set is approximately 1.0 mL
Slowly infuse lidocaine 40 mg IO over 120 seconds
Allow lidocaine to dwell in IO space 60 seconds
Flush with 5-10 mL of normal saline
Slowly administer an additional 20 mg of lidocaine IO over 60 seconds
Repeat PRN
Consider systemic pain control for patients not responding to IO lidocaine
Recommended anesthetic for infant/child responsive to pain:
Observe recommended cautions/contraindications to using 2% preservative and epinephrine free lidocaine (intravenous lidocaine)
Confirm lidocaine dose per institutional protocol; usual initial dose is 0.5 mg/kg, not to exceed 40 mg
Prime extension set with lidocaine Note that the priming volume of the EZ-Connect® Extension Set is approximately 1.0 mL
  For small doses of lidocaine, consider administering by carefully attaching syringe directly to needle hub (prime extension set with normal saline)
Slowly infuse lidocaine over 120 seconds
  Allow lidocaine to dwell in IO space 60 seconds
Flush with 2-5 mL of normal saline
Slowly administer subsequent lidocaine (half the initial dose) IO over 60 seconds
  Repeat PRN
Consider systemic pain control for patients not responding to IO lidocaine
The use of any medication, including lidocaine, given IV or IO is the responsibility of the treating physician, medical director or qualified prescriber and not an official recommendation of Teleflex Incorporated and its subsidiaries. Teleflex is not the manufacturer of lidocaine, and the user should be familiar with the manufacturer’s instructions or directions for use for all indications, side-effects, contraindications, precautions and warnings of lidocaine. Teleflex disclaims all liability for the use, application or interpretation of the use of this information in the medical treatment of any patient. Lidocaine dosing recommendations were developed based on research; for additional information, please visit www.eziocomfort.com

## Slide 27: Assessment/Documentation
Assess frequently to ensure:
    Wrist band is on patient
    No evidence of complications
      Catheter is intact/patent
      No evidence of infiltration/ extravasation
    Dressing & connections are secure
    Repeat flush as needed
    Need for initial or additional lidocaine
Removal
Notes: Care and maintenance of the IO catheter is similar to other venous access routes. Ensure the site is intact and catheter is securely in place, connections are secure and dressing is clean, dry & intact.
Confirm catheter patency/placement prior to and throughout medication and fluid administration.
Repeat a rapid normal saline flush as needed
Monitor frequently for evidence of complications.
Infiltration/extravasation is the most common complication associated with IO insertion, and can lead to serious complications such as compartment syndrome and necrosis. The catheter insertion site and posterior compartments (depending on insertion site) should be monitored frequently for any signs of infiltration/extravasation, localized inflammation, limb perfusion or dislodgement. As with any line, manipulation and patient transport are critical times for ensuring patency remains intact.
Assess patient comfort and manage infusion related pain.
Consider adding an IO documentation template to your computer charting.
The catheter may remain in place up to 24 hours and is not MRI compatible.

## Slide 28: ARROW® EZ-IO® Catheter Removal
Notes: Attach a luer lock syringe to the hub of the catheter. Rotate syringe and catheter clockwise while pulling straight out to withdraw catheter. Do not rock or bend during removal to prevent the hub breaking off the catheter. Dispose of all sharps and biohazard materials using standard biohazard practices and disposal containers.
Dress site as appropriate. There are no activity restrictions after the catheter is removed.
The patient should be directed to seek medical care if they experience signs and symptoms of infection, pain near the insertion site or any other unusual symptoms.

## Slide 29: EZ-IO® Driver
Clean per your institution’s protocol or Instructions for Use
  Check to ensure nothing has attached to the magnetic tip
Inspect driver and return to case or replace trigger guard
Do NOT submerge or autoclave driver!
Notes: Drivers are sealed and not intended to be opened. Batteries are not replaceable.
The driver is manufactured using a non-porous, fluid-resistant, medical-grade plastic covering; however, it is not fluid-proof. For cleaning purposes, the driver should never be submerged in any cleaning solution. Follow the driver’s Instructions for Use when cleaning. Check to ensure that nothing has attached to the magnetic tip. If NOT storing driver in case, the trigger guard may be placed to secure the trigger from accidental discharge/battery waste. The red battery indicator light will illuminate when approximately 10% of battery life remains.
If your clinical environment requires sterilization, the driver can be sterilized using the STERRAD® 100S, NX Standard cycle, and 100NX Standard cycle.
Refer to the Instructions For Use (IFU) for more detailed information.

## Slide 30: Connect With Us
24 Hour Clinical Support 800-680-4911
Education Resources
www.teleflex.com/ezioeducation
Connect with the App
for iPhone & Android
30
Notes: 24/7 live clinical support is available by calling 800-680-4911. This number can be found on our patient wristband and our website. We place tremendous value on direct clinical feedback, so we invite you to visit our website to gain the newest, cutting edge product information as well as provide us with your direct feedback on our product’s performance.
We know how busy you are, and we are committed to providing you with the resources necessary to ensure policy development, education and competency is a seamless, comprehensive process. Our experienced, professional clinicians are available to provide inservices, speaking engagements and professional education programs. Additionally, these and other clinical professional resources are available:
Detailed educational videos appropriate for health care providers and patients
FAQs booklet: The Science & Fundamentals of Intraosseous Vascular Access
Policy and procedure template
Hand outs, educational brochures and posters
Online guided presentation with examination and completion certificate
Educational PowerPoint presentations
Instructional videos
Competency checklist
Written quiz
QI/QA form
Difficult Vascular Access (DVA) algorithm
And many other resources – please let us know what you are looking for!

## Slide 31: THANK YOU
STERRAD is a product of Advanced Sterilization Products, a Johnson and Johnson Company.
Robertsite is a product by Halkey-Roberts Corporation.
NeedleVISE is a product of Atrion Medical Products, Inc.
Level 1 is a product of the Medical Division of Smiths Group, plc.
Teleflex, EZ-IO, EZ-Connect and EZ-Stabilizer are trademarks or registered trademarks of Teleflex Incorporated or its affiliates.
Vidacare LLC is a wholly owned subsidiary of Teleflex Incorporated.
© 2014 Teleflex Incorporated. All rights reserved.
This material is not intended to replace standard clinical education and training by Vidacare LLC, a subsidiary of Teleflex Incorporated, and should be utilized as an adjunct to more detailed information which is available about the proper use of the product. View educational resources at www.teleflex.com/ezioeducation or contact a Teleflex clinical professional for any detailed questions related to product insertion, maintenance, removal and other clinical education information.