EZ-IO® Proximal Humerus Insertion Site Identification – Infant/Small Child

- Place the patient’s hand over the abdomen (elbow adducted and humerus internally rotated)

- Place your palm on the patient’s shoulder anteriorly

  - The area that feels like a “ball” under your palm is the general target area

  - You should be able to feel this ball, even on obese patients, by pushing deeply

- Place the ulnar aspect of your hand vertically over the axilla

- Place the ulnar aspect of your other hand along the midline of the upper arm laterally

- Place your thumbs together over the arm

  - This identifies the vertical line of insertion on the proximal humerus

- Palpate deeply up the humerus to the surgical neck

  - This may feel like a golf ball on a tee – the spot where the “ball” meets the “tee” is the surgical neck

  - The insertion site is 1 to 2cm above the surgical neck, on the most prominent aspect of the greater tubercle

EZ-IO® Proximal Humerus Insertion Technique – Infant/Small Child

- Prepare the site by using antiseptic of your choice

- Use a clean, “no touch” technique

- Remove the needle set cap

- Point the needle set tip at a 45-degree angle to the anterior plane and posteromedial

- Push the needle set tip through the skin until the tip rests against the bone

- The 5mm mark must be visible above the skin for confirmation of adequate needle set length

- Gently drill, immediately release the trigger when you feel the loss of resistance as the needle set enters the medullary space

  - Avoid recoil – do NOT pull back on the driver when releasing the trigger

- Hold the hub in place and pull the driver straight off

- Continue to hold the hub while twisting the stylet off the hub with counter clockwise rotations

  - The catheter should feel firmly seated in the bone (1st confirmation of placement)

- Place the stylet in a sharps container

- Place the EZ-StabilizerTM dressing over the hub

- Attach a primed EZ- Connect®  extension set to the hub, firmly secure by twisting clockwise

- Pull the tabs off the EZ-Stabilizer dressing to expose the adhesive, apply to the skin

- Aspirate for blood/bone marrow (2nd confirmation of placement)

- Secure the arm in place across the abdomen

Recommended Anesthetic for Infant/Child Responsive to Pain:

- Observe recommended cautions/contraindications to using 2% preservative and epinephrine free lidocaine (intravenous lidocaine)

- Confirm lidocaine dose per institutional protocol

Usual initial dose is 0.5mg/kg, not to exceed 40mg

- Prime EZ-Connect extension set with lidocaine

Note that the priming volume of the EZ-Connect is approximately 1.0mL

For small doses of lidocaine, consider administering by carefully attaching syringe directly to needle hub (prime EZ-Connect with normal saline)

- Slowly infuse lidocaine over 120 seconds

Allow lidocaine to dwell in IO space 60 seconds

- Flush with 2-5 mL of normal saline

- Slowly administer subsequent lidocaine (half the initial dose) IO over 60 seconds

Repeat PRN

- Consider systemic pain control for patients not responding to IO lidocaine

Infant/Child Unresponsive to Pain

- Prime EZ-Connect extension set with normal saline

- Flush the IO catheter with 2-5 mL of normal saline

- Connect fluids if ordered, infusion may need to be pressurized to achieve desired rate

- Assess for any signs of extravasation/complications

Should patient develop signs that indicate responsiveness to pain, refer to section “Recommended anesthetic for infant/child responsive to pain”

EZ-IO® Removal Technique

- Remove EZ-Connect and EZ-Stabilizer dressing

- Stabilize catheter hub and attach a Luer lock syringe to the hub

- Maintaining axial alignment, twist clockwise and pull straight out

  - Do not rock the syringe

- Dispose of catheter with syringe attached into sharps container

Lidocaine dosing recommendations were developed based on the research below.  For additional references, research and dosing charts, please visit www.eziocomfort.com

- Philbeck TE, Miller LJ, Montez D, Puga T. Hurts so good; easing IO pain and pressure. JEMS 2010;35(9):58-69*

- Ong MEH, Chan YH, Oh JJ, Ngo AS-Y. An observational, prospective study comparing tibial and humeral intraosseous access using the EZ-IO. Am J Emerg Med 2009;27:8-15*

- Fowler RL, Pierce A, Nazeer S et al. 1,199 case series: Powered intraosseous insertion provides safe and effective vascular access for emergency patients. Ann Emerg Med 2008;52:S152*

- Paxton JH, Knuth TE, Klausner HA. Proximal humerus intraosseous infusion: a preferred emergency venous access. J Trauma. 2009; 67: 606-11*

- Wayne MA. Intraosseous vascular access: devices, sites and rationale for IO use. JEMS 2007;32:S23-5.

- Frascone RJ, Jensen JP, Kaye K, Salzman JG. Consecutive field trials using two different intraosseous devices. Prehosp Emerg Care 2007;11:164-71*

- Fowler R, Gallagher JV, Isaacs SM, et al. The role of intraosseous vascular access in the out-of-hospital environment (resource document to NAEMSP position statement). Prehosp Emerg Care 2007;11:63-6

- Miller L, Kramer GC, Bolleter S. Rescue access made easy. JEMS 2005;30:S8-18*

- Davidoff J, Fowler R, Gordon D, et al. Clinical evaluation of a novel intraosseous device for adults: prospective, 250-patient, multi-center trial. JEMS 2005;30:S20-3*

- Gillum L, Kovar J. Powered intraosseous access in the prehospital setting: MCHD EMS puts the EZ-IO to the test. JEMS 2005;30:S24-6*

- Cooper BR, Mahoney PF, Hodgetts TJ, Mellor A. Intra-osseous access (EZIO®) for resuscitation: UK military combat experience. JR Army Med Corps 2008;153(4):314-6.

- Hixson R. Intraosseous administration of preservative-free lidocaine. http://www.vidacare.com/files/Hixson-Lidocaine-%20032012.pdf. Accessed November 22, 2013.

*Research sponsored by Vidacare Corporation

Vidacare disclaims all liability for the use, application or interpretation of this information in the medical treatment of any patient.